Agilent 1200 Series HPLC — Dry Run (Clickthrough Only)

Goal: prime pumps, isocratic flow with water, watch pressure trace, save method  ·  No column required

SUPPLIES (minimal)

| ☐ | Bottle of HPLC-grade water (small) |
| --- | --- |
| ☐ | Empty waste bottle if not plumbed |
| ☐ | USB stick |

For dry run, you can run with column REMOVED (replace with a union) to avoid column equilibration time. Otherwise budget +20–30 min for column equilibration.

POWER UP

| ☐ | Power on stack: degasser → quaternary pump → autosampler → column compartment → DAD (left-to-right or bottom-to-top) |
| --- | --- |
| ☐ | Wait ~3–5 min for all modules to initialize |
| ☐ | Launch ChemStation; verify all modules show "Online / Ready" |

PRIME / PURGE

| ☐ | Confirm solvent A bottle (HPLC water) is filled and inlet line submerged |
| --- | --- |
| ☐ | Open the purge valve on the pump (front panel) |
| ☐ | Pump → Purge → channel A → 5 mL/min for 3 min |
| ☐ | Close purge valve |
| ☐ | Set isocratic flow: 100% A, 0.5 mL/min |

SOFTWARE CLICKTHROUGH

| ☐ | Method → New Method → isocratic 100% water, 0.5 mL/min, 5 min total |
| --- | --- |
| ☐ | DAD: signal A 254 nm, ref 360 nm, range 200–400 nm scan |
| ☐ | Sample info: name = "dryrun_blank" |
| ☐ | Click Run Method (no injection — just flow) |
| ☐ | Watch pressure trace stabilize (with no column: very low, ~5–20 bar) |
| ☐ | Watch DAD signal: flat baseline expected |
| ☐ | When done, File → Save Method, Save Data → export to USB |

SHUTDOWN

| ☐ | Reduce flow to 0 mL/min; pump off |
| --- | --- |
| ☐ | DAD lamp off (ChemStation → DAD → Lamp Off) |
| ☐ | Close ChemStation; leave stack powered |
| ☐ | Pack out |

TIMING & BOTTLENECK LOG

| Step | Time | Notes / bottleneck |
| --- | --- | --- |
| Power up + initialize |  |  |
| Software online |  |  |
| Prime / purge pump |  |  |
| Isocratic flow (~5 min) |  |  |
| Save method + data |  |  |
| Shutdown |  |  |
| TOTAL on-site time |  |  |

Bottlenecks / supplies that ran short / things to add to next Amazon order: