Bruker Tensor 27 + Hyperion FT-IR Microscope — Dry Run

Goal: power up → OPUS login → background scan → "sample" scan on empty stage → save

SUPPLIES (minimal)

| ☐ | USB stick for saving the .0 spectrum file |
| --- | --- |
| ☐ | (Optional) one 13 mm KBr disc for a more representative background — empty stage is fine for clickthrough |
| ☐ | Lint-free wipes if you handle the KBr disc |

OPUS 5.5 default password is OPUS (all caps). KBr/NaCl discs source via Amazon, not the Chemstore.

POWER UP

| ☐ | Power on Tensor 27 (rear rocker); confirm green status LED on front |
| --- | --- |
| ☐ | If using Hyperion microscope path: power on Hyperion controller separately |
| ☐ | Wait 15–30 min for IR source thermal stability (you can clickthrough during warmup) |
| ☐ | Power on PC; launch OPUS 5.5; login user OPUS, password OPUS |

SOFTWARE CLICKTHROUGH — TENSOR 27 (bench)

| ☐ | OPUS → Measure → Advanced Measurement |
| --- | --- |
| ☐ | Experiment: load default Tensor method (or Optics: Setup → confirm DTGS or MCT detector + KBr beamsplitter) |
| ☐ | Resolution 4 cm⁻¹, scan range 4000–400 cm⁻¹, 16 background scans, 16 sample scans |
| ☐ | Sample form: empty stage (sample compartment open) |
| ☐ | Click "Background Single Channel" — wait for completion (~30 sec) |
| ☐ | Click "Sample Single Channel" with stage still empty — should give a near-100% transmission trace |
| ☐ | File → Save File As → .0 to USB |

SOFTWARE CLICKTHROUGH — HYPERION (microscope, optional)

| ☐ | Switch optical path to microscope (front-panel slider or OPUS Optics dialog) |
| --- | --- |
| ☐ | Open Hyperion video window in OPUS |
| ☐ | Move stage joystick to a clear/empty area on the sample holder |
| ☐ | Focus using the trinocular and condenser |
| ☐ | Set 15× Cassegrain objective, aperture ~50×50 µm |
| ☐ | Click Background → Sample on the same empty area (sample = background → flat trace, this is correct) |
| ☐ | Save .0 file to USB |

If MCT detector is in use, top up with liquid N2 and let stabilize 30 min before measuring — not needed for clickthrough on DTGS.

SHUTDOWN

| ☐ | Close OPUS |
| --- | --- |
| ☐ | Switch optical path back to bench (if you used Hyperion) |
| ☐ | Leave Tensor 27 powered (lab convention — source stability) |
| ☐ | Power off Hyperion controller if separately powered |
| ☐ | Pack out |

TIMING & BOTTLENECK LOG

| Step | Time | Notes / bottleneck |
| --- | --- | --- |
| Power up + IR source warm |  |  |
| OPUS launch + login |  |  |
| Tensor: BG + sample scan |  |  |
| Hyperion: stage + scan (opt) |  |  |
| Save .0 file(s) |  |  |
| Shutdown + pack out |  |  |
| TOTAL on-site time |  |  |

Bottlenecks / supplies that ran short / things to add to next Amazon order: