Horiba Jobin Yvon FluoroMax-3 — Dry Run (Clickthrough Only)

Steady-state spectrofluorometer  ·  150 W xenon lamp  ·  Goal: emission scan with empty cuvette

SUPPLIES (minimal)

| ☐ | 4-clear-side fluorescence quartz cuvette (1 cm) — EMPTY is fine |
| --- | --- |
| ☐ | Kimwipes + nitrile gloves |
| ☐ | USB stick |

Fluorescence cuvettes have all four faces polished — do not use the 2-clear-face UV cuvettes used for absorbance.

POWER UP

| ☐ | Power on the lamp housing (xenon arc lamp ignites with audible click) |
| --- | --- |
| ☐ | Wait 30 min for lamp output to stabilize (you can clickthrough during) |
| ☐ | Power on PC and launch FluorEssence (or DataMax/Origin if older install) |
| ☐ | Verify FluoroMax-3 module shows "Online" |

SOFTWARE CLICKTHROUGH (FluorEssence)

| ☐ | Experiment Menu → Spectra → Emission Scan |
| --- | --- |
| ☐ | Excitation wavelength: 350 nm · Emission range: 360 → 600 nm |
| ☐ | Slits: 5 nm excitation, 5 nm emission · Integration 0.1 s · Step 1 nm |
| ☐ | Detector: S1 (signal) ratioed to R1 (reference) |
| ☐ | Place empty cuvette in sample holder; close lid |
| ☐ | Click Run — should complete in ~60 sec |
| ☐ | Confirm spectrum window: expect Raman scatter peak from air ~395 nm, otherwise flat |
| ☐ | File → Save → .spc or .txt to USB |

If detector saturates ("signal too high"), close excitation slit to 2 nm and rerun. Never operate with the lid open — the PMT is sensitive to room light.

SHUTDOWN

| ☐ | Close FluorEssence |
| --- | --- |
| ☐ | Turn off xenon lamp (cool-down ~10 min before re-igniting) |
| ☐ | Leave PC and module powered |
| ☐ | Pack out |

TIMING & BOTTLENECK LOG

| Step | Time | Notes / bottleneck |
| --- | --- | --- |
| Lamp on → stable |  |  |
| Software launch + online |  |  |
| Method setup |  |  |
| Empty-cell emission scan |  |  |
| Save spectrum |  |  |
| Lamp cool + shutdown |  |  |
| TOTAL on-site time |  |  |

Bottlenecks / supplies that ran short / things to add to next Amazon order: