Jasco J-1500-150 CD — Dry Run (Clickthrough Only)

Room CB-016A  ·  Goal: walk the full software path with empty cuvette, no real sample

PRE-RUN (start ~30 min before)

| ☐ | Confirm Apex Nevis 10C is connected to J-1500 purge inlet (1/4" BSPP) |
| --- | --- |
| ☐ | Power Nevis 10C: front rocker to "1", then breakers right → left → middle |
| ☐ | Wait for blue N2 LED on generator (means N2 is flowing, even if reservoir not yet at 99.9995%) |
| ☐ | Open J-1500 purge valve; let N2 flow into spectrometer ~20 min |

SUPPLIES (minimal)

| ☐ | 1 cm or 1 mm quartz cuvette — EMPTY is fine for clickthrough |
| --- | --- |
| ☐ | Kimwipes + nitrile gloves |
| ☐ | USB stick for saving the test .jws file |

No standards, no solvents, no buffer. This run is only validating the software path.

POWER UP J-1500

| ☐ | Power on J-1500 main unit (rear switch) |
| --- | --- |
| ☐ | Start xenon lamp via Spectra Manager II OR front-panel button |
| ☐ | Wait 15 min for lamp to stabilize (you can clickthrough menus during this time) |

SOFTWARE CLICKTHROUGH (Spectra Manager II)

| ☐ | Login as default user; verify the J-1500 module shows "Online / Ready" |
| --- | --- |
| ☐ | Open: Spectrum Measurement |
| ☐ | File → New (or load default CD method) |
| ☐ | Set Parameters → Range: 400 → 260 nm (avoids ozone risk and O2 absorption) |
| ☐ | Bandwidth: 2 nm · Scan speed: 200 nm/min · D.I.T.: 0.5 s · Accumulations: 1 |
| ☐ | Channels: CD + HT (high voltage) + ABS |
| ☐ | Place empty cuvette in sample compartment, close lid |
| ☐ | Click "Baseline" — should complete in ~40 sec |
| ☐ | Click "Sample" — will produce a flat trace (this is correct, no sample present) |
| ☐ | File → Save As → .jws file to USB |
| ☐ | File → Export → ASCII (.txt) to USB |

If Spectra Manager II throws "PMT overload / HT > 600V", drop a 10 mm cell of water in to attenuate.

SHUTDOWN

| ☐ | Close Spectra Manager II |
| --- | --- |
| ☐ | Turn off xenon lamp |
| ☐ | Power off J-1500 main unit |
| ☐ | LEAVE Nevis 10C running in standby — do NOT switch off the rocker (avoids 5-hr re-purge) |
| ☐ | Pack out (no storage allowed) |

TIMING & BOTTLENECK LOG

| Step | Time | Notes / bottleneck |
| --- | --- | --- |
| Generator on → blue N2 LED |  |  |
| Spec compartment purge (~20 min) |  |  |
| Lamp on → stable |  |  |
| Software clickthrough |  |  |
| Empty-cell baseline + scan |  |  |
| Save + export to USB |  |  |
| Shutdown + pack out |  |  |
| TOTAL on-site time |  |  |

Bottlenecks / supplies that ran short / things to add to next Amazon order: