Shimadzu MALDI-8020 MALDI-TOF — Dry Run (Clickthrough Only)

Benchtop linear MALDI-TOF  ·  MALDIsolutions software  ·  Goal: load empty plate, fire laser at blank spot, save

SUPPLIES (minimal)

| ☐ | Clean MALDI target plate (any standard Shimadzu plate) |
| --- | --- |
| ☐ | Lint-free wipes + methanol (for plate clean) |
| ☐ | Nitrile gloves |
| ☐ | USB stick |

MALDI-8020 has built-in vacuum pump — no separate rough/turbo to babysit. Pumpdown is much faster than AXIMA (~3–5 min vs 15–20).

POWER UP

| ☐ | Power on the MALDI-8020 (rear switch); wait for status LED steady (~2 min) |
| --- | --- |
| ☐ | Launch MALDIsolutions on PC; verify instrument shows "Ready" |
| ☐ | Confirm laser interlock: chamber lid closed, status "Safe" |

LOAD EMPTY TARGET PLATE

| ☐ | Software → Sample → Vent (waits ~30–60 sec to atmosphere) |
| --- | --- |
| ☐ | Open chamber drawer, slide in clean target plate, close drawer |
| ☐ | Software → Sample → Pump Down (~3–5 min to operating vacuum) |

SOFTWARE CLICKTHROUGH (MALDIsolutions)

| ☐ | Open Method Editor; load default Linear Positive method |
| --- | --- |
| ☐ | Confirm laser power LOW (e.g., 30/180 setting) and 50 shots, 1 profile |
| ☐ | Mass range 500–3000 m/z |
| ☐ | Open Acquisition window; in camera view pick a clean spot on the empty plate |
| ☐ | Click Acquire — laser fires, spectrum window populates with flat baseline (correct, no analyte) |
| ☐ | Verify the data file path and detector trace are visible |
| ☐ | File → Save → .lcd / .run file (depends on install) to USB |

If you see real peaks, plate isn't clean — not blocking for the dry run, just note it.

SHUTDOWN

| ☐ | Vent chamber, remove plate, store in plate holder |
| --- | --- |
| ☐ | Pump back down (instrument prefers being stored under vacuum) |
| ☐ | Close MALDIsolutions |
| ☐ | Leave instrument powered (vacuum maintenance) |
| ☐ | Pack out |

TIMING & BOTTLENECK LOG

| Step | Time | Notes / bottleneck |
| --- | --- | --- |
| Power up + software ready |  |  |
| Vent + load empty plate |  |  |
| Pump down (~3–5 min) |  |  |
| Method setup clickthrough |  |  |
| Laser fire on blank |  |  |
| Save data file |  |  |
| Vent + repump for storage |  |  |
| TOTAL on-site time |  |  |

Bottlenecks / supplies that ran short / things to add to next Amazon order: