TA Instruments DSC Q100 — Dry Run (Clickthrough Only)

Goal: load empty Tzero pans, run a quick 25→60°C ramp, confirm data save  ·  No sample  ·  Older sibling of Q20

SUPPLIES (minimal)

| ☐ | 2 empty aluminum Tzero pans + lids |
| --- | --- |
| ☐ | Tweezers |
| ☐ | USB stick |

POWER UP

| ☐ | Confirm N2 cell purge ~50 mL/min |
| --- | --- |
| ☐ | Power on Q100 module, wait for ready (~5 min) |
| ☐ | Power on RCS or LNCS cooler if equipped |
| ☐ | Launch Universal Analysis 4 (or TA Instrument Explorer if older install) and verify Q100 is online |

Q100 software stack is older than Q20 — some labs run Universal Analysis 4 alongside Advantage.

LOAD EMPTY PANS

| ☐ | Open cell lifter |
| --- | --- |
| ☐ | Empty pan + lid on SAMPLE thermocouple |
| ☐ | Empty pan + lid on REFERENCE thermocouple |
| ☐ | Lower inner and outer cell lids |

SOFTWARE CLICKTHROUGH

| ☐ | New Experiment → Mode: Standard DSC |
| --- | --- |
| ☐ | Sample: name = "dryrun", mass = 0.001 mg |
| ☐ | Method: Equilibrate 25°C → Ramp 10°C/min to 60°C → End |
| ☐ | Apply → Start |
| ☐ | Watch heat flow trace: nearly flat baseline expected |
| ☐ | Save → .001 file to USB |

SHUTDOWN

| ☐ | Wait for cell <40°C, open lid, remove pans |
| --- | --- |
| ☐ | Power off cooler if applicable; leave Q100 powered |
| ☐ | Close software, pack out |

TIMING & BOTTLENECK LOG

| Step | Time | Notes / bottleneck |
| --- | --- | --- |
| Power up + software online |  |  |
| N2 purge stable |  |  |
| Load empty pans |  |  |
| Method setup clickthrough |  |  |
| Run (ramp 25→60°C) |  |  |
| Save .001 file |  |  |
| Cool + unload |  |  |
| TOTAL on-site time |  |  |

Bottlenecks / supplies that ran short / things to add to next Amazon order: