TA Instruments DSC Q20 — Dry Run (Clickthrough Only)

Goal: load empty Tzero pans, run a quick 30→50°C ramp, confirm data save  ·  No sample

SUPPLIES (minimal)

| ☐ | 2 empty aluminum Tzero pans + lids (one sample, one reference) |
| --- | --- |
| ☐ | Tweezers (the lab tweezers should be on the bench) |
| ☐ | USB stick for saving .001 file |

No sample, no encapsulation press needed for this run.

POWER UP

| ☐ | Confirm N2 purge gas connected: cell purge ~50 mL/min |
| --- | --- |
| ☐ | Power on the Q20 module (rear switch); wait for green ready LED (~3–5 min) |
| ☐ | Power on RCS90 cooler if equipped (otherwise ambient run only) |
| ☐ | Launch TA Instrument Explorer (or Trios) on the PC; verify Q20 is online |

LOAD EMPTY PANS

| ☐ | Open the cell lifter (silver button, top of module) |
| --- | --- |
| ☐ | Place empty pan + crimped lid on the SAMPLE thermocouple (rear position by convention) |
| ☐ | Place empty pan + crimped lid on the REFERENCE thermocouple (front) |
| ☐ | Lower the inner and outer cell lids slowly |

SOFTWARE CLICKTHROUGH

| ☐ | Open Q Series Advantage → Experiment View |
| --- | --- |
| ☐ | New Experiment → Mode: Standard |
| ☐ | Sample info: name = "dryrun", mass = 0.001 mg (placeholder — nonzero required) |
| ☐ | Method: Equilibrate 30°C → Ramp 10°C/min to 50°C → End |
| ☐ | Notes: "clickthrough dry run, empty pans" |
| ☐ | Click "Apply" then "Start" (green arrow) |
| ☐ | Watch heat flow trace: should be nearly flat baseline |
| ☐ | When run ends (~3 min), click "Save" and write .001 file to USB |

If RCS is on, you can re-equilibrate to 30°C; otherwise let cell drift to ambient.

SHUTDOWN

| ☐ | Wait for cell to return to <40°C before opening lid |
| --- | --- |
| ☐ | Lift lid, remove both pans with tweezers, dispose |
| ☐ | Close lid |
| ☐ | Power off RCS90 if applicable |
| ☐ | Leave Q20 module powered (instrument prefers continuous power); close software |
| ☐ | Pack out |

TIMING & BOTTLENECK LOG

| Step | Time | Notes / bottleneck |
| --- | --- | --- |
| Power up + software online |  |  |
| N2 purge stable |  |  |
| Load empty pans |  |  |
| Method setup clickthrough |  |  |
| Run (ramp 30→50°C) |  |  |
| Save .001 file |  |  |
| Cool down + unload |  |  |
| TOTAL on-site time |  |  |

Bottlenecks / supplies that ran short / things to add to next Amazon order: