VTSYIQI Digital Visible Spectrophotometer — Dry Run (Clickthrough Only)

Portable visible-range spec  ·  Goal: power up, blank-calibrate, take blank reading, log

SUPPLIES (minimal)

| ☐ | Plastic or glass cuvette (10 mm) — the unit ships with a holder for these |
| --- | --- |
| ☐ | Distilled or DI water for blank |
| ☐ | Kimwipes |
| ☐ | Notebook (display-only readout, no PC export) |

Confirm the unit is fully charged or has a working power adapter before walking in — portable units sometimes don't survive transit storage.

POWER UP

| ☐ | Power on; wait for splash + auto-self-check (~30 sec) |
| --- | --- |
| ☐ | Tungsten lamp warms up (~10 min for stable readings) |
| ☐ | Set wavelength using up/down keys: e.g., 540 nm (visible green) |

BLANK CALIBRATION

| ☐ | Fill cuvette with DI water; wipe outside dry; insert into sample holder |
| --- | --- |
| ☐ | Press "0 ABS / 100% T" button to calibrate baseline |
| ☐ | Display should read 0.000 ABS or 100.0% T |

READING

| ☐ | Leave the same water-blank cuvette in place |
| --- | --- |
| ☐ | Press "Measure" or "Read" |
| ☐ | Confirm reading within ~±0.005 ABS of zero |
| ☐ | Step wavelength: try 400, 540, 700 nm; recalibrate at each, take blank read |
| ☐ | Log values in notebook |

SHUTDOWN

| ☐ | Remove cuvette, dry and store |
| --- | --- |
| ☐ | Power off |
| ☐ | Pack out |

TIMING & BOTTLENECK LOG

| Step | Time | Notes / bottleneck |
| --- | --- | --- |
| Power up + lamp warm |  |  |
| Blank calibration |  |  |
| Reading at λ1 |  |  |
| Re-cal + reading at λ2 |  |  |
| Re-cal + reading at λ3 |  |  |
| Shutdown |  |  |
| TOTAL on-site time |  |  |

Bottlenecks / supplies that ran short / things to add to next Amazon order: