TA Instruments TGA Q50 — Dry Run (Clickthrough Only)

Goal: tare empty Pt pan, run a quick 30→100°C ramp, confirm weight signal & data save

SUPPLIES (minimal)

| ☐ | Empty platinum sample pan (lab usually keeps a tray of clean pans by the instrument) |
| --- | --- |
| ☐ | Tweezers / pan hook |
| ☐ | USB stick for saving .001 file |

Pt pans are reusable — no need to bring new ones for the dry run.

POWER UP

| ☐ | Confirm N2 cylinder open: balance purge ~40 mL/min, furnace purge ~60 mL/min |
| --- | --- |
| ☐ | Power on Q50 module; wait for green ready LED (~5 min) |
| ☐ | Launch Q Series Advantage on PC; verify Q50 is online |
| ☐ | Confirm balance reads stable (typically displays current pan weight or 0) |

LOAD EMPTY PAN + TARE

| ☐ | If autosampler is equipped: place clean Pt pan in autosampler tray position 1 |
| --- | --- |
| ☐ | OR manually: open furnace, hang empty Pt pan on hangdown wire with tweezers |
| ☐ | Software → Control → Tare (zeroes the balance with empty pan) |
| ☐ | Wait for tare to complete (~30 sec); confirm display reads ~0.000 mg |

SOFTWARE CLICKTHROUGH

| ☐ | Experiment View → New |
| --- | --- |
| ☐ | Sample: name = "dryrun_empty", mass = 0.001 mg (placeholder) |
| ☐ | Method: Equilibrate 30°C → Ramp 20°C/min to 100°C → End |
| ☐ | Gas: N2 throughout |
| ☐ | Notes: "clickthrough, empty Pt pan" |
| ☐ | Click Apply → Start |
| ☐ | Furnace closes, ramp begins; confirm weight trace stays near 0 (±0.01 mg drift OK) |
| ☐ | When complete (~4 min), Save → .001 file to USB |

Drift > 0.05 mg over the ramp suggests the balance needs more warmup or there is contamination on the pan.

SHUTDOWN

| ☐ | Wait for furnace to cool to <50°C before opening |
| --- | --- |
| ☐ | Remove Pt pan, return to clean pan tray |
| ☐ | Close furnace, close software |
| ☐ | Leave Q50 powered + N2 flowing at low rate (lab convention) |
| ☐ | Pack out |

TIMING & BOTTLENECK LOG

| Step | Time | Notes / bottleneck |
| --- | --- | --- |
| Power up + N2 stable |  |  |
| Load + tare empty pan |  |  |
| Method setup clickthrough |  |  |
| Run (ramp 30→100°C) |  |  |
| Save .001 file |  |  |
| Furnace cool + unload |  |  |
| TOTAL on-site time |  |  |

Bottlenecks / supplies that ran short / things to add to next Amazon order: