Agilent 7890 GC-FID (Chiral) — Dry Run (Clickthrough Only)

Goal: light the FID, run a blank temperature program, save  ·  No injection

SUPPLIES (minimal)

| ☐ | USB stick for data export |
| --- | --- |
| ☐ | (Optional) clean empty syringe in autosampler |

PRE-VISIT GAS CHECK

| ☐ | Carrier (He or H2) cylinder open, regulated to ~50 psi |
| --- | --- |
| ☐ | FID H2 cylinder open, regulated to ~30 psi |
| ☐ | FID air supply on (house air or cylinder), ~50 psi |
| ☐ | Vent hood / exhaust active above the GC |

POWER UP / LIGHT FID

| ☐ | Power on the 7890 GC; wait for ready (~3 min) |
| --- | --- |
| ☐ | Launch GC ChemStation; verify Instrument 1 online |
| ☐ | Method → load any chiral column method (or default short method) |
| ☐ | Detector → FID → turn on H2 + Air flows; wait ~30 sec |
| ☐ | Detector → FID → Light Flame (igniter clicks; signal jumps then settles) |
| ☐ | Confirm baseline signal stable at ~5–20 pA |

If FID won't light, check H2 flow > 30 mL/min and air flow ~400 mL/min. Re-try ignition.

BLANK RUN

| ☐ | Method: oven program 50°C hold 1 min → 100°C @ 30°C/min, hold 1 min |
| --- | --- |
| ☐ | Sample info: "dryrun_blank", no injection |
| ☐ | Click Run; total ~5 min |
| ☐ | Watch FID trace: flat baseline, slight thermal drift OK |
| ☐ | Save → .D folder → export to USB |

SHUTDOWN

| ☐ | Detector → FID Off (extinguishes flame, closes H2) |
| --- | --- |
| ☐ | Return oven to standby (~50°C) |
| ☐ | Close ChemStation; leave PC + GC powered |
| ☐ | Pack out |

TIMING & BOTTLENECK LOG

| Step | Time | Notes / bottleneck |
| --- | --- | --- |
| Gas check + power up |  |  |
| Software online |  |  |
| FID light |  |  |
| Blank run (~5 min) |  |  |
| Save + export |  |  |
| FID off + standby |  |  |
| TOTAL on-site time |  |  |

Bottlenecks / supplies that ran short / things to add to next Amazon order: