Agilent 7890A GC + 5975C Inert MSD — Dry Run (Clickthrough Only)

Goal: confirm vacuum, autotune MSD, run blank temperature program, save  ·  No injection

SUPPLIES (minimal)

| ☐ | USB stick for data export |
| --- | --- |
| ☐ | PFTBA tune compound is internal to the MSD — no consumables needed |

If MSD has been vented, pumpdown to operating vacuum from cold can take 1–2 hours. Verify it's already pumped before walking in.

PRE-VISIT CHECK

| ☐ | Confirm carrier gas (He) cylinder pressure adequate (>500 psi) |
| --- | --- |
| ☐ | Confirm MSD foreline pump is running (audible) and turbo at speed |
| ☐ | Confirm heated zones are at standby temps (transfer line ~280°C, source ~230°C) |

POWER UP / SOFTWARE

| ☐ | Power on PC if off; launch MSD ChemStation → Instrument 1 (Online) |
| --- | --- |
| ☐ | Wait for instrument communication ("Ready" status) |
| ☐ | View → Tune and Vacuum Control |
| ☐ | Confirm vacuum: ion gauge < 2×10⁻⁵ Torr |

AUTOTUNE

| ☐ | Tune → Autotune ("atune.u" or "PFTBA tune") |
| --- | --- |
| ☐ | Wait ~5–10 min for tune to complete |
| ☐ | Inspect tune report: confirm 69, 219, 502 m/z peaks present and abundances within spec |
| ☐ | Save tune report |

BLANK RUN

| ☐ | Method → Load Method → select short test method (e.g., 50°C hold 1 min → 200°C @ 30°C/min, hold 1 min) |
| --- | --- |
| ☐ | Sequence not needed for single run — use Run Method directly |
| ☐ | Sample info: name = "dryrun_blank" |
| ☐ | Click Run — NO injection (or insert empty syringe in autosampler) |
| ☐ | Watch TIC trace (~6 min total): expect flat baseline with possible solvent ghost |
| ☐ | When complete, File → Save Data; Export to .CDF or .D folder to USB |

SHUTDOWN

| ☐ | Return GC to standby method (low oven temp) |
| --- | --- |
| ☐ | Close ChemStation; leave PC + MSD pumping |
| ☐ | Pack out |

TIMING & BOTTLENECK LOG

| Step | Time | Notes / bottleneck |
| --- | --- | --- |
| Power up + ChemStation online |  |  |
| Vacuum check |  |  |
| Autotune (~5–10 min) |  |  |
| Method load + setup |  |  |
| Blank run (~6 min) |  |  |
| Save data + export |  |  |
| TOTAL on-site time |  |  |

Bottlenecks / supplies that ran short / things to add to next Amazon order: