Jandel RM3 Four-Point Probe — Dry Run (Clickthrough Only)

Goal: power up, lower probe head, take a reading on a test substrate, log

SUPPLIES (minimal)

| ☐ | A test substrate — a clean piece of conductive material (e.g., aluminum foil flat-mounted, or a copper sheet) |
| --- | --- |
| ☐ | Notebook (RM3 doesn't always export to USB — readings may need to be hand-logged) |
| ☐ | Tweezers |

Probe tips are spring-loaded tungsten carbide and easily damaged — lower the head SLOWLY and never onto a hard non-flat surface.

POWER UP

| ☐ | Power on the RM3 controller (rear switch) |
| --- | --- |
| ☐ | Confirm display shows current/voltage range selectors |
| ☐ | Verify probe head is in raised position before placing sample |

MEASUREMENT CLICKTHROUGH

| ☐ | Place test substrate flat on stage under probe head |
| --- | --- |
| ☐ | Lower probe head SLOWLY using the manual lift until tips contact the surface (gentle spring compression) |
| ☐ | Select current range appropriate for material (e.g., 1 mA for metals, 100 µA for semiconductors) |
| ☐ | Press Measure |
| ☐ | Read sheet resistance (Ω/sq) or resistivity from display |
| ☐ | Log reading in notebook |
| ☐ | Raise probe head; reposition; take a second reading at a new spot |
| ☐ | Confirm readings reasonable (foil ~mOhm/sq, copper similar) |

SHUTDOWN

| ☐ | Raise probe head fully |
| --- | --- |
| ☐ | Remove substrate |
| ☐ | Power off |
| ☐ | Pack out (do not store probe with head down) |

TIMING & BOTTLENECK LOG

| Step | Time | Notes / bottleneck |
| --- | --- | --- |
| Power up |  |  |
| Sample placement |  |  |
| Probe lower + measure |  |  |
| Second-spot measurement |  |  |
| Log readings |  |  |
| Shutdown |  |  |
| TOTAL on-site time |  |  |

Bottlenecks / supplies that ran short / things to add to next Amazon order: